## List of Items
v1-11-Mar-2022
Item description | unit | Total required quantity | Minimum qty per donation (per item)
Priority medicines
AMOXICILLIN + CLAVULANIC ACID, 1000 mg + 200 mg powder for inj | vial | 610000 | 121000
CEFOTAXIME, 2g powder for inj | vial | 80000 | 14400
CLINDAMYCIN,150mg/ ml powder for inj , | amp | 90000 | 17300
CLOXACILLIN 500 mg, powder for inj.(I.V.) | vial | 540000 | 107600
ENOXAPARIN, 0.4 ml | prefiled syringe | 80000 | 14400
METOPROLOL TARTRATE INJECTION, USP 5MG/5ML (1MG PER ML) AMP | amp | 10000 | 2000
NOREPINEPHRINE TARTRATE | amp | 30000 | 4800
ONDANSETRON, 2 mg/ml sol inj | amp | 100000 | 19200
DEXAMETHASONE, 4mg/ml 2ml | amp | 140000 | 26900
FENTANYL, 50mcg/ml in 2 ml | amp | 30000 | 4800
SERUM ANTIGANGRENOUS, Equine polyvalent, pruified, 30,000 IU per dose | amp | 30000 | 2000
POTASSIUM IODINE, tablets, 125mg | tablet | 2000000 | 250000
Equipments
PATIENT MONITOR,w/CO2 IBP ECG NIBP TEMP, SPO2, w/acc. | unit | 300 | 10
ANAESTHESIA MACHINE, closed circuit, trolley, w/acc | unit | 15 | 3
ECG RECORDER,portable, 12 leads, with printer and access | unit | 300 | 10
DEFIBRILLATOR,AED,w/access | unit | 1000 | 50
SYRINGE PUMP, single channel, AC and battery powered,w/acc. | unit | 300 | 10
INFUSION PUMP, LCD, flow 0.1-1500mL/h, 220V, batt., w/acc | unit | 300 | 10
EXAMINATION LIGHT, Led, overhead, mobile, ajustable, on wheels | unit | 240 | 20
PUMP,SUCTION,FOOT-OPERATED, with tubes and connector | unit | 1000 | 50
PATIENT VENTILATOR, intensive care, for adult and pediatric, with breathing circuits and patient interface (TYPE 2): this category of ventilators is recommended for use in contexts where oxygen supply disruptions may occur. These models feature an oxygen inlet that can be temporarily used with a lower pressure oxygen source | unit | 15 | 3
SCANNER,ULTRASOUND,mobile, w/access | unit | 240 | 20
SELF-INFLATING BAG SET, self-refilling, capacity > 1500 m adult + 3 masks (S, M, L) | unit | 240 | 40
CAPNOMETER portable, EtCO2/RR monitoring, AAA battery, w/acc. | unit | 240 | 20
X-RAY UNIT BASIC, mobile | unit | 15 | 3
AMBULANCE, van type or 4WD, LHD, ALS or ACLS | unit | 20 | 2
SURGICAL DRILL, cordless +accessories + drill bit | unit | 50 | 5
DERMATOME, ELECTRICAL with battery + access. | unit | 50 | 5
LEG TRACTION SPLINT, pre Hospital care and transport | unit | 1000 | 50
Consumables
CENTRAL VENOUS CATHETERS, Y and straight needle , triple lumen | unit | 1000 | 50
SET, INTRAOSSEOUS INFUSION KIT, w/ needles and IV set | set | 750 | 20
SET, INFUSION, adult, sterile, s.u. | set | 800000 | 50000
SET, INFUSION, pediatric, precision, sterile, s.u. | set | 200000 | 20000
DRAIN, THORACIC, INSERTION-SET, complet, sterile, disposabl | set | 2000 | 200
Surgical instruments and fixators (please see tab "Set of Surgery instrument" with detailled content)
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, DRESSING | set | 2000 | 50
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | set | 240 | 10
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | set | 240 | 10
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | set | 240 | 10
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | set | 240 | 10
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | set | 240 | 10
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | set | 240 | 10
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | set | 240 | 10
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | set | 240 | 10
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | set | 240 | 10
SET, EXTERNAL FIXATION, LARGE,FIXATORS & INSTRU | set | 1180 | 10
SET, EXTERNAL FIXATION, LARGE,FIXATORS & INSTRU | set | 500 | 10

## Technical Specifications
PATIENT MONITOR,w/CO2 IBP ECG NIBP TEMP, SPO2, w/acc. | Continuous display on screen of patient ECG, respiration rate and heart rate, invasive and non-invasive blood pressure, body temperature, SpO2 and CO2 mainstream or side stream. Dynamic digital display that can show all active parameters. Unwanted parameters can be deselected from display. Operator can set audio-visual alarm levels for low or high levels of each parameter independently. Operates from mains voltage or from internal rechargeable battery. ECG patient connectors that are sterilizable and reusable are preferred. Hard copy printout of traces will not be required. • Multichannel (up to 12 leads) ECG measurement and selectable display; an extra option for simple five-lead connection preferred. • Heart rate measurement range to be at least 30–250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm. • SpO2 measurement range at least 70–99%, with accuracy better than ± 3% and minimum gradation 1%. • Blood pressure monitoring range at least 30–270 mmHg, minimum gradation 1 mmHg. Cuff sizes neonatal/paediatric and adult. User selectable measurement intervals. • Internal pump for cuff inflation for non-invasive blood pressure measurement, with over pressure protection. • Temperature probe to be reusable, external skin contact type, but consumable to protect between patients or disinfection method explained. • Temperature range at least 30–40 °C, minimum gradation 0.1 °C. • Respiration rate measurement range at least 0–100 bpm, minimum gradation 1 bpm. • CO2 monitoring capabilities. • Invasive blood pressure (IBP) monitoring capabilities. Automatic and programmable memory. Storage of continuous monitoring data. Trace signal velocity of at least 25 mm/s. LCD or TFT screen with: • analogue shape signals and numerical values visualization; • settable limits for the measured variables; • not less than 10”wide. Design must enable use in demanding environments, e.g. shock, vibration and free fall tests as per tests. Protections of all the functions against defibrillator discharges and electrosurgical units. Pace-maker detection. Data management functions (preferable). Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85% (90% preferable). Enclosure to have ingress protection level IPX1 or better.
ANAESTHESIA MACHINE, closed circuit, trolley, w/acc | Suitable for adult, paediatric and newborn patients. The unit is mounted on a trolley with minimum of four (4) anti-static swivel castors. A minimum of two of the castors are provided with breaks. The unit is equipped with an upper shelf. The cart has been proved with handles for manoeuvring. A side rail for mounting accessories is included. The unit has an adjustable patient-circuit support arm. The unit is provided with a minimum of three (3) gas inlets (O₂, N₂O and Air). Gas inlet connections are compliant with DISS or NIST (needs to be specified when ordering). The unit is provided with a minimum of two (2) gas cylinder yokes (O₂, and N₂O). Holders for bottles are to include a secure locking mechanism. The minimum gas supply input pressure range should be between 2.8 – 6 bar (41 - 87 psi). The unit are equipped with gas supply gauges with scales allowing easy readout. All pipeline connections have diameter-indexed safety systems (DISSs), or another means of preventing connection of dangerous gases. The unit is equipped with a minimum of four (4) analogue rotameters, two (2) for O₂, one (1) for air and one (1) for N₂O with programmed parameters visualization. The O₂ and N₂O flowmeters have a minimum range of 0.0 – 10 L/min, and a resolution of at least 0.2 L/min. Externally supplied N₂O, O₂, and air mixture ratios are fully controllable. There are separate controls for the flow of: O₂, N₂0, and anaesthetic agents. Flows and the mixture ratios determined from flowmeter settings are accurate to within ± 10% of set values or better. The unit has the ability to carry out self-diagnosis and integrity testing, including a compliance and leakage test. The unit is equipped with a minimum of two (2) vaporizer slots with a Selectatec mounting system. The vaporizers are provided with an interlock system, preventing the use of more than one vaporizer simultaneously. One (1) Halothane vaporizer is included with the unit. The unit basic is equipped with a passive scavenging system and optional allow for an active scavenging system. The unit is equipped with a soda lime CO₂ absorber. The unit is equipped with a non-return and a three-way valve, including the connection tube. A built-in rechargeable battery allows for an autonomy of at least one (1) hour. The unit will switch back and forth between battery and mains operation in case of power failures. Recharging resumes automatically when extern power becomes available. The unit is designed and constructed in such a way that it is able to withstand frequent disinfection with hospital-grade products. Fuses are fitted on both live and neutral supply lines. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit).
GLUCOMETER, with 300 strip & lancet | Linear Range: 0-20 mmol/L (0-360 mg/dL) Results Format: Quantitative across whole linear range (should be able to switch between mg and mmol) Result Units: mg/dL OR mmol/L Precision: ±2% or 2.5 mg/dL, whichever is greater Sample: Whole blood heel-stick sample <50 μL Requiring no calibration Kit Stability & Storage: le for 12 months with harsh ambient conditions (temperature 10-40 °C, humidity 15%-95% elevation up to 2000 meters) and transport stress (48h with fluctuations up to 50°C and down to 0°C)
ECG RECORDER,portable, 12 leads, with printer and access | Automatic equipment calibration. 12 electrographic standard leads measuring with simultaneous acquisition. Visualization of at least one group of 3 leads simultaneously. At least 1 rest ECG patient cable with at least 5 terminals. LCD or TFT colour monitor display of at least 7 inches with visualization of analogical curves, alphanumeric values measured and the related physiological limits. Frequency response range at least: 0.05–150 Hz. Equipment compatible with patients with pacemakers. Protection against defibrillation. Automatic internal data storing up to at least 40 ECG records. Input impedance not less than 5 MΩ (preferably higher than 10 MΩ). Patient leakage current not higher than 10 uA. Common mode rejection (CMR) not less than + 100 dB (preferably + 120 dB). ECG signal measurement range not smaller than -2 mV to +2 mV. At least the following adjustable alarms: a) heart failure b) ventricular fibrillation c) tachycardia d) bradycardia e) electrode disconnection. Equipment provided with filters at least for: baseline instability, AC 50/60 Hz interference, rumours, low- pass, etc. Equipment provided with at least the following software applications: a) arrhythmias detection and analysis b) ventricular fibrillation detection and analysis c) ventricular tachycardia detection and analysis d) ST segment analysis. Integrated printer using standard A4 format paper. Selectable printing paper speed at least of 5, 25 and 50 mm/sec. Displayed parameters Arrhythmia detection basic. Numeric display. Graphical and numerical trends and physiological curves display (preferable).
Otoscope (ear) | Specifications: Viewing lens magnification 2.5-3x Color temperature: Cool white in the range 3100-5500K Light intensity: 8,000-12,000 lux Specula specifications Reusable Autoclavable Cleanable with alcohol wipes Handle with on/off switch. Scratch-resistant lens; glass (preferred) or plastic Port for an insufflation bulb Set contained in a hard storage case. Liquid splash resistant Weight: < 0.170kg without batteries Batteries: Standard AA or AAA or rechargeable Battery life: at least 3 hours of use on full charge or fresh batteries Low battery or charge status indicator (as applicable to rechargeable type only) Supplied without batteries (as applicable to AA/AAA type only) Storage: Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
ophthalmoscope (eye) | Product Description: Ophthalmoscope set to visualize the interior of the eye, with the instrument relatively close to the subject's eye and the observer viewing an upright magnified image. Ophthalmoscope set composed of diagnostic head threaded on a handle. Specifications Range of lenses not smaller than -20D to +20D with steps not greater than 1D. Anti-reflection lens. Magnification 12-16x Apertures: small, large and semi circle, fixation star. Color temperature: Cool white in the range 3100-5500K Light intensity: 8,000-12,000 lux Dust free sealed optics and aspherical optical system Red-free, blue, green, and polarization filters. Handle with on/off switch. Scratchproof lens; glass (preferred) or plastic Batteries: AA or rechargeable Battery life: at least 5 hours of use on full charge or fresh batteries Set contained in a hard storage case. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Liquid splash resistant Weight: < 0.20kg without batteries Supplied without batteries
DEFIBRILLATOR,AED,w/access | TECHNICAL SPECIFICATIONS Adult and paediatric settings. Biphasic output waveform Automatic ECG (VT/VF) detection and analysis. Automatic impedance compensation based on patient. Paediatric dose attenuation. Automatic switch between AED and CPR modes based on analysis. Analysis time less than 10 seconds after having been switched on. Maximum energy output for adults between 150 and 200 Joules. Maximum energy output for paediatrics 50 Joules, or adjustable between 30 and 70 (depending on the model supplied). Charge time to maximum energy output is 5 seconds. Includes step-by-step device and CPR user guide, either in durable plastic-coated manual and/or on machine. Audible metronome for CPR procedure. Built-in discharge feature for safety. Integrated control panel with all parameters and controls. Step-by-step pictograms on the control panel for ready and easy operation. Audio and/or visual indications of operational status and step-by-step operation. Audio and/or visual alarms for operational status, electrodes, battery status and system errors. Automatic self-test and continuous check of pads and electrodes connection. Conductive surface area of adult electrodes is at least 80 cm² Conductive surface area of paediatric electrodes is at least 80 cm² Shelf life of electrodes is at least 2 years. Replaceable internal battery, non-rechargeable. Battery type LiSO2 or LiM no (depending on the model supplied). The internal battery, when full supports at least 140 full discharges at 200J, or two hours continuous ECG monitoring. Weighs less than 3 kg. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x carry case, with storage pockets for leads and accessories. 3 x set of adult adhesive external pads, colour-coded and with pictograms. 3 x set of paediatric adhesive external pads, colour-coded and with pictograms. 1 x plastic enclosed Quick Reference Guide (step-by-step AED and CPR). 1 x CD-ROM/DVD/USB with training material. 2 x sets of spare batteries, packed separately.
SYRINGE PUMP, single channel, AC and battery powered,w/acc. | TECHNICAL SPECIFICATIONS Open system, compatible with all standard brands of syringes. Compatible with 10, 20, 30 and 50mL capacity syringes at a minimum. Continuous delivery, linear motor and plunger driven. The unit is equipped with occlusion detection. Self-test is performed each time the device is switched on. User programmable for syringe size, infusion volume, time, and flow rate. Automatic calculation of third parameter when user enters in other two (volume, time, and flow rate). Minimum guaranteed flow rate of 0.1-1300mL/hr, depending on syringe size. Minimum Keep Vein Open (KVO) rate of 0.1-1mL/hr but never greater than programmed flow rate. Volume delivered with an accuracy of at least 3%. Maximum pressure of at least 17.4 PSI / 120 kPa. User adjustable high pressure/occlusion settings. Display includes start/stop, volume limit, flow rate and volume so far delivered. System reports with audio-visual alert on operational status such as, but not limited to ready, end-of-injection. System reports with audio-visual alert on malfunctions such as, but not limited to syringe position, occlusion, low/high flow. System reports with audio-visual alert on low battery status. Ability to silence audio alarms for maximum of 2 minutes. Built in battery with a capacity to run the unit for 7 hours at 5mL/hr flow rate. Automatic switch from mains to battery in case of power failure. Automatic battery charge when mains connection is re-established. Capable of being mounted on mobile pole/(roll) stand, bed rail and wall-mounted rails. Designed for frequent and easy dismount and disinfection with hospital-grade products. Power requirements: 240 Volts - 50/60 Hz (110 Volts can be supplied if specifically requested).
INFUSION PUMP, LCD, flow 0.1-1500mL/h, 220V, batt., w/acc | One channel (at least). Capable of accepting any kind of fluids (solutions and medications preferable). Pump capabilities: • flow range 0.1 to ≥ 999 mL/hr • increments 0.1–100 mL/hr • keep vein open (KVO) rate 1–5 mL/hr • volume to be infused selector (VTBI) 1–9999 mL • flow rate accuracy of ± 5% or better • when multiple channel automatic piggybacking. Ingress protection not less than IPX2. Front panel lockout. Self-check carried out on powering on. Events stored system – log book. Pause infusion facility required. Anti-bolus system to reduce pressure on sudden release of occlusion. IV set: • free-flow protection • air trapping capability • needleless IV connection. “Dose error” reduction system (preferable). Drug library software available, including updates (free during warranty). Air bubble detector with single and cumulative functions (preferable). Clearly visible optical alarms. Acoustic alarms not less than 45dB. Easy set up and operation. Large easy to read display. Real-time display. Availability of a nurse call system connectable to a staff alerting system, 24 V/0.2 A; static or dynamic (preferable). Compatibility with standard infusion sets commonly distributed in the market (desirable at least by the leading brands). Equipment provided “ready to use”. Continuous operation within specification in ambient temperature of at least 5–40 °C, relative humidity of at least 10–90% non-condensing. A limit of 85% relative humidity will be possibly considered as “acceptable” in the current COVID-19 pandemic context. Any accessory or dedicated device necessary to the proper functioning and utilization of the equipment is included. Shock-fall protection.
EXAMINATION LIGHT, Led, overhead, mobile, ajustable, on wheels | TECHNICAL SPECIFICATIONS Heavy base with low centre of gravity. Single light head with LED lights in cluster. Articulating arm between light and base stand. Base with 5 anti-static swivel castors in a star formation. At least two castors have been equipped with brakes. Minimum 14 cm fixed focus field of view. Minimum illumination of 60,000 lux at 0.5m. Minimum colour temperature between 4,000°K. Minimum colour rendering index (CRI) of 95. Minimum LED life more than 40,000 hours. Floor to light height adjustable to include at least range of 1.1 m to 1.75 m. On/off switch. Power requirements: 100 - 240 Volts - 50/60 Hz. SUPPLIED WITH Instructions for assembly, use and maintenance in English. OPTIONAL ACCESSORIES AVAILABLE - An integrated rechargeable battery with a minimum of 2 hours battery life can be supplied upon request at extra cost.
PUMP,SUCTION,FOOT-OPERATED, with tubes and connector | Oil free vacuum pump. Maximum vacuum not less than 450 mmHg (adjustable by control). Maximum suction capacity not less than 15 L/min. Collection bottle (1 or 2, preferably 2): at least 1 L capacity each bottle (disposable bag or collection jar), preferably 2 L. Bottle(s) to have an automatic cut off when full to prevent ingress of fluid to pump. Filter and overflow valve incorporated to prevent cross-contamination (e.g. shatterproof material, over flow protection system). It should be disposable or autoclavable. Equipment control: manual on/off power switch and, preferably also foot switch. Air line to pump to incorporate bacterial filter. Tubing to patient to be minimum 1.5 m long, non-collapsible type. All parts are manufactured from high-strength, durable material, that does not require specific maintenance or storage conditions. Pump can be disassembled entirely, is easy to clean, disinfect and sterilize. Equipment provided complete with dedicated trolley/cart or housing with castor wheels for easy movement. Sound level not higher than 60 dBA. System integrated holder for suction cannulas/tubing easy and safe positioning. Any necessary greasing/oiling to be simple, accessible and possible by health users.
PATIENT VENTILATOR, intensive care, for adult and pediatric, with breathing circuits and patient interface (TYPE 2) | PATIENT VENTILATOR, intensive care, for adult and pediatric, with breathing circuits and patient interface (TYPE 2): this category of ventilators is recommended for use in contexts where oxygen supply disruptions may occur. These models feature an oxygen inlet that can be temporarily used with a lower pressure oxygen source Medical oxygen and air high-pressure input ports (> 35 psi [2.4bar]) provide a means to limit reverse gas flowrate (leakage). Each high-pressure input port with a filter and water trap, if applicable, for air input port. Medical air compressor or turbine in-built preferred, alternatively external air compressor. Possibility for using external low-pressure oxygen (approx. 20 psi), as source (preferable). Mechanical safety valve. Internal function testing/leak testing. At least IP21 degree of protection to the harmful ingress of water (fluid spill resistance). Mechanical shock resistance, mechanical vibration, electromagnetic compatibility and electrical safety testing. Compatible active humidifying system. Event log for errors traceability (preferable). Operating temperature and humidity 5–40 °C and 0–95% relative humidity (RH). Storage temperature and humidity -20–60 °C, 0–95% RH. Ventilation modes Pressure control ventilation (PCV). Volume control ventilation (VCV). Pressure support ventilation (PSV). Synchronized intermittent mandatory ventilation (SIMV) (preferable). Pressure regulated volume control (PRVC) or similar (preferable). Non-invasive ventilation (CPAP or BiPAP).
SCANNER,ULTRASOUND,mobile, w/access | TECHNICAL SPECIFICATIONS General purpose ultrasound scanner, with obstetrics package and probes. System includes scanner and printer mounted on a trolley. The system supports B mode and M mode. The unit includes: one (1) convex array probe, 2.5-5 MHz and one (1) endovaginal probe, 5.5-7.5 MHz. The unit supports user programmable exams (exam pre-sets). An obstetrics analysis package is included which provides for foetal weight and anatomy, gestational age and foetal growth calculations. The unit comes standard with packages for: abdomen, gynaecology/obstetrics, cardiology and small parts. Depending on the model supplied Pulsed Wave Doppler (back/white), is included in the unit. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). IMAGE DISPLAY Digital callipers. Cine loop. A minimum of 4 x zoom. 256 levels of greyscale. A freeze function. Capable of lateral and vertical inversion in B mode. Capable of displaying B+M modes simultaneously. Capable of displaying 2B and 4B modes. Full screen annotation, for real-time and frozen images. COMPONENTS Time gain control (TGC). Tissue harmonic imaging. Optional B-mode optimization. The unit has a range of pre- and post-processing fetures to enhance images and automated image optimization. A black and white display with a minimum of 25 cm (9.7 inch) diagonally. Equipped with an alphanumeric keyboard and a trackball. Two probe ports can be connected simultaneously, machine selectable from interface. PRINTER A digital printer is included with the unit. Data connection to ultrasound. support 256 greyscale. 325 dpi resolution at a minimum. Includes printer head cleaning kit. CART/TROLLEY Includes at least two side probe holders. Side handle(s) for transport and positioning. Surface space for scanner and keyboard. Storage space (shelf) for printer. Four anti-static braking swivel castors, at least 10 cm high. DATA STORAGE AND COMMUNICATIONS Internal hard drive for limited data storage. Supported communication protocols: USB, S-video, VGA, and depending on the model: RS 232 or ethernet). Supported formats for image export: JPEG. PERFORMANCE AND SAFETY Self-test is performed each time the machine is switched on. System reports, with audio-visual alert on: operational status and malfunctions.
Laryngoscope with blades of different sizes | Product specifications: Large hollow, cylindrical, slightly ribbed handle Handle made of either chromium-plated or stainless steel Can be opened to insert two batteries (type LR14, size C, 1.5 V) Stud contact, fitting various sizes and types of depressors With a set of four stainless steel depressors, with halogen bulb Maclntosh type: Curved Nr 2, length approx. 110 mm Curved Nr 3, length approx. 135 mm Curved Nr 4, length approx. 155 mm Miller type: Straight Nr 1, length approx. 100 mm Supplied with: 1 x Durable protective plastic box or padded vinyl case 4 x spare halogen bulb (one for each depressor) 1 x Instructions for assembly, use and maintenance in English Required, but not supplied: 2 x S1802211 Dry cell, alkaline, "C", 1.5V.
SELF-INFLATING BAG SET, self-refilling, capacity > 1500 m adult + 3 masks (S, M, L) | Hand-operated resuscitator used for mechanical ventilation of adult and paediatric patients. Easy to disassemble and reassemble. Easy to clean and disinfect. Reusable. All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions. Ventilation can be done with ambient air or with oxygen. Resuscitator shall be supplied as a complete set with: • Compressible self-refilling ventilation bag with maximum capacity not less than 1300 mL. Dead volume: < 7 mL. • One-way valve with or without pressure limiting. Pressure-limiting system: the elasticity of the bag’s outer cover limits the airway pressure to approximately 7 kPa (70 cmH2O) when squeezed normally with one hand. • Patient connector, outside diameter: 22 mm; inside diameter: 15 mm. • Inlet valve with nipple for oxygen tubing. • Oxygen reservoir bag capacity according to patient size.
CAPNOMETER portable, EtCO2/RR monitoring, AAA battery, w/acc. | CO2 detection device for verification of correct endotracheal tube placement; single use. Usage time: up to 2 hours preferable. Connector port adapted to standard patient circuit (22 Fr/15 mm). Colour chart for different ranges of EtCO2. Sizes Adult: internal volume: 25 cc approximately. Paediatric: internal volume: 3 cc approximately.
X-RAY UNIT BASIC, mobile | Equipment easy to assemble/install/set up and be fully functional and operational as a complete standalone solution for acquisition, review, presentation, display, storage and transfer of radiographic images in a resource-limited setting. Equipment: mobile, motorized, battery and AC power operated (AC mandatory for charging and, preferably, for standard working operations too). Digital Imaging and Communications in Medicine (DICOM) 3.0 compatible image storage and transfer required. The system should be capable of storing at least 2000 images with comprehensive post-processing options (Dose Area Product [DAP] meter to be integrated). Capacity for removable media storage, to transfer data through different options (CD, DVD and/or USB), to send images through existing network port, and, preferably, to have wireless transfer of images through hospital wireless network (wireless and cable connections shall be provided). Integrated Ethernet connectivity required. At least 20 anatomical programmes shall be available. Software for COVID-19 detection should be included (preferable) (if available). An integrated image review monitor to be included in the configuration. Anti-scatter grid or software for scatter correction (preferable) (if available). Equipment provided with DAP device/capability to record the patient dose.
LEG TRACTION SPLINT, pre Hospital care and transport | Mechanical tubular frame including foot plate 2 leg supports: upper and lower part of the leg Pneumatic pump (possible manual adjustment) Adjustable length: extension from 64 to114 cm Pressure gauge allowing traction control, and possible adjustment during transport Fixation with Velcro straps at foot heigth and securing bridle with button fastener at knee and thigh height Recommended traction strength: 4.5 to 18 kg Security air release valve
SET, EXTERNAL FIXATION, LARGE,FIXATORS & INSTRU | The external fixator set which is proposed in the surgical instruments sets allows to face most of the situations as well as in adults as in children. Each box allows 8 to 10 assemblies.. The selected equipment is easy to use, particularly by non orthopaedic surgeons. The use of external fixator sets implies precise and regular management of their composition. All items can be ordered separately. The main indications of external fixators are open fractures. Therapeutic alternatives such as orthopaedic treatment by wound debridement-traction, fenestrated plastering or amputation should be considered. Remote training for clinical procedures.
SURGICAL DRILL, cordless +accessories | Operating Cycle: Forward / back running: cycle of 5 times 1 min running/ 4 min off , with 2 h break Operating in oscillation mode: cycle of 5 times 15 sec running /15 sec off, with 90 min break Operating with sagittal saw tip: cycle of 4 times 10 sec running / 20 sec off, with 90 min break Speed: Drill Position: 1500 tr/min, approx ReamerPosition : 303 tr/min, approx Couple: Drill Position: 3,3 Nm, approx. Reamer Position: 15,7 Nm, approx. Motor cannulated: 4,0 mm, approx. Double trigger with forward, backward, alternating mode Internal power supply: 9,9 V (DC) battery pack (lithium-ion) 9,6 V per battery pack (nickel-cadmium) Battery life: 20 min under continuous use with Lithium-Ion Water Ingress: At least IPX6 Dimensions H x W x L:137 x 34 x 103 mm. approx Weight Handpiece: 0,58 kg, approx Standard accessories: Straight, angled, angled double-lock, curved bur, straight variable exposure, angled variable exposure, footed, rotating footed, contra-angle, metal cutting, telescoping, angled telescoping, taper telescoping, perforator driver, Jacobs, chuck, bone mill
DERMATOME, ELECTRICAL with battery + access. | Blade travel: 3.1 mm 6300–6800 strokes/min Width of cut: 78 mm maximum adjustable with 10 flaps which are 8 mm wide Thickness of cut: 0.2 mm to 1.2 mm graduation every 0.1 mm Dimensions: 285 x 115 x 50 mm Autoclavable à 134ºC (= 2 bars) or 121ºC (= 1 bar)
Ambulance, EN 1789 Type B or C (mobile intensive care units) | Advanced ambulance passing ECS Regulatory Authority and in road-worthy state: a. Lights, windscreen wipers, brakes, seatbelts, doors, and warning devices must be functional. b. The windscreen must not impair driver vision. c. Tires must have appropriate tread and not be overly worn as to be dangerous for use. d. License/identification plate must be clear and visible per governmental and/or regional standards. e. A working fire extinguisher should be available at all times. f. Fuel levels should be adequately supplied for the expected use of the vehicle. g. Oil, brake fluid, radiator water, and battery should be maintained per local public safety vehicle ordinance standards. h. All ambulances should be insured according to local standards (for example for collision and liability coverage). i. Appropriate and functioning personnel and patient restraints, including seatbelts, door locks, and a stretcher lock. Equipped with: IV infusion flow regulator (manual with dial) Intravenous infusion set (lines, connectors and cannulas size 14 - 24) IV poles or hooks Tourniquets for IV start Intraosseous needle driver (electric or manual) Intraosseous needle or equivalent (15, 25 and 45mm) Fluid warmer Pressure bag for IV infusion Rapid infusion catheter set Automated external defibrillator (AED) Syringe (2, 5, 10, 20, 50cc) Needles (range of gauges and lengths; safety, single use) Oral medication administration supplies (eg. cups, dropper) Lubricating jelly Patient tags (eg. bracelets) Cotton wool Nasogastric (NG) tube (5 - 18F) Urinals Bed pans Kidney basins Trauma shears Sterile procedure drapes Non-sterile procedure drapes Non-sterile pad with impermeable barrier Razor Scalpel Adhesive tape Stocked high acuity grab bag Sterile dressings Vaseline or paraffin gauze Paediatric length-based tape (e.g. Broselow tape) Child weighing scale (250g graduation) Adult weighing scale Clock, timer or watch with second hand Stethoscope Digital thermometer (32 - 43 Celsius) Penlight Pulse oximetry Blood pressure (BP) cuff (small - large adult) Paediatric sizes for blood pressure cuffs Tongue depressors Glucometer test strips Quality controls for glucometer Finger-stick lancets Glucometer Plaster cast remover Plaster of paris or equivalent Sheet or binder for pelvic fractures Splint material (i.e. - stocking, gauze padding, premade splint) Tension or traction splint (e.g. hare splint) Elastic bandages Cervical collars (small - large) Paediatric sizes for cervical collars Arterial tourniquet Delivery set Cord clamp Birth kit Neonatal resuscitation kit Sanitary pads Suction catheters Suction collection vessel Suction device: manual (bulb) or foot pump Suction tubing (#10 - 16) and connectors Yankauer or other stiff suction tip Oxygen supply (cylinder, concentrator or other source; with flow meter) Nasal prongs (adult, paediatric and neonatal sizes) Oxygen mask (adult, paediatric and neonatal sizes) Non-rebreather Oxygen tubing and connectors Oral airways (#00 - 6) Nasal airways (12 - 36F) Bag–valve–mask (adult, paediatric and neonatal sizes) PEEP valve for bag-valve-masks Metered-dose inhaler spacer Paediatric sizes for laryngoscope Endotracheal tubes (#2.5 - 8.5) Bougie Surgical cricothyroidotomy set Laryngoscope set (range of blades and sizes) Oesophageal detector device (e.g. CO2 colorimeter) Dual-tube laryngeal mask airway (LMA) (#2 - 7) Physical Restraints Sterile dressing set Non-sterile dressings System to record care and provide transition instructions Cards for communication number distribution Important telephone number/radio frequency list Condition specific protocols (posters or cards/book) Triage protocols (posters and cards) Stretchers and gurneys (wheeled stretchers) Linens (e.g. sheets, pillow cases, towels, patient gowns) Pillows Headlamps Lockable box for controlled medications Medication and medication dispensing storage and organization system Supplies and equipment storage and organization system (eg. shelving, drawers, cart) Shortwave radio or dedicated telephone line Patient slide for stretcher transfers Goggles Rescue blankets Pill crusher Tablet cutter Fire extinguishing mechanism (e.g. blankets, extinguishers) Soap Hand disinfectant Environmental disinfectant Cleaning equipment (e.g. sponges, buckets, mops, brooms) Goggles Mask with face shield Surgical mask N-95 respirators Non-sterile gowns (multiple sizes) Sterile gloves (size 6 - 8) Safe biological waste disposal containers Red and yellow plastic bags Rubbish bags Sharps disposal Ethanol 70% solution Safe final disposal of biological waste Safe final disposal of sharps Printed book of disaster protocols
SET, INFUSION, adult, sterile, s.u. | General Description: Infusion giving set, sterile, single use Technical Specifications: Sterile single use device used for parenteral administration of infusions (IV fluids) by gravity. Device components: Perforator: Sharp piercing device with protective cap. Air inlet: With integrated bacteriological filter. Drop counting chamber: Transparent drip chamber, calibrated at 20 drops/ml, with 15-20μm fluid filter. Tubing: Transparent tubing, minimum length 150cm, with injection site latex or latex-free (or Y-injection port), with distal connector preferably Luer Lock connector. Precision flow regulator: Smooth roller clamp to facilitate easy, safe control and adjustment of fluid rates. Sterile Single use Conform to ISO 8536-4 Infusion equipment for medical use
SET, INFUSION, pediatric, precision, sterile, s.u. | Technical Specifications: Sterile IV giving set with graduated chamber (burette) enabling precise volume and slow administration of infusion or injectable medicine. Tube: plastic (PVC: polyvinyl chloride). Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is approximately 3mm / 4mm. Plastic Perforator: plastic (polyacetate). Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Graduated chamber (burette): plastic (PS: polystyrene). Transparent. Capacity: approximately 100-150ml. Graduations every ml. Connected to the perforator by a tube (approximatey 15cm). Security gauge: Automatic shut off valve, to stop the infusion as soon as the pre-set volume has flown away in order to prevent air entering the drop-counting chamber and/or the tube. Drop-counting chamber: Located under the graduated chamber. Calibrated 60 drops/ml. Fitted with a 15µ polyamide filter. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Single use. Sterile.
SET, INTRAOSSEOUS INFUSION KIT, w/ needles and IV set | BATTERY POWERED MEDICAL DRILL hand-held waterproof power driver powererd by sealed lithium battery single use battery that cannot be replaced. The battery indicator (LED) will blink red when the trigger is activated and has only 10% of battery life remaining. life expectancy: 10 years or +/- 500 insertions no maintenance required dimensions : approximately 16.5 cm x 11.4 cm x 6.4 cm weight: less than 0.5 kg NEEDLE + STABILIZER SET 5 x IO needles, 15G (Ø 1.8 mm) austenitic stainless steel needle (18/8) markings / 5 mm indicating depth, 5 mm of the catheter (at least one black line) must remain visible Luer lock connector with stylet for insertion and protective cap sterile, for single use size: paediatric: 15 mm long, pink, for patients weighing between 3 to 39kg adult: 25 mm long, blue, for patients weighing more than 39kg 5 x extension set adapted to IO needle needed to connect the IO needle safely to the infusion set ​with clamp sterile, for single use 5 x security wristband to document date and time of insertion 5 x ADHESIVE DRESSING sterile, single use size : 14.6 cm x 13.3 cm x 3.2 cm latex-free 5 x NeedleVISE 1 port: device for securing the needle after insertion, to be disposed of in a sharps container. TRANSPORT BAG for IO kit
CENTRAL VENOUS CATHETERS, Y and straight needle , double lumen (triple preferred) | double, or triple preferred, lumen central venous catheter with extension tubes and clamps outer diameter: CH 7 for adults, CH 5 for children, CH4 for newborns. soft catheter tip X-ray detectable, opaque polyurethane length: CH7: 20 cm for left insertion, 16 cm for right insertion, CH4: 8 cm, CH5: 11 cm. length markings for accurate catheter placement no antibacterial components included in the catheter material 1x16G and 2x18G kink-proof guide wire (stainless steel or nitinol) with advancer and flexible J-tip dilator puncture needle needle- free connector per lumen. surgical scalpel fixation wing for securing the catheter

## Set of Surgery instrument
Set decpetion | Set content | Qty per set
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | SCISSORS, NELSON-METZENBAUM, 23 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | PHOTOMACROGRAPHIC SCALE, 100mm, , black aluminum, PPSM300 | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | KIDNEY DISH, small, lg. 170 mm, , stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | SCALPEL, HANDLE, No 3 (for blades 10/11/15) | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | SCALPEL, HANDLE, No 4 (for blades 20) | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | SCISSORS, MAYO, 17 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | SCISSORS, METZENBAUM, 18 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, DRESSING, STANDARD, 20 cm, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, TISSUE, STANDARD, straight, 2-3 teeth, 20 cm | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | NEEDLE HOLDER, MAYO-HEGAR, 15cm, standard | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, HEMOSTATIC, H-MOSQUITO, 12.5 cm, curved | 6
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, HEMOSTATIC, CRILE, 14cm, straight | 6
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 6
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, HEMOSTATIC, R-OCHSNER, 16 cm/1x2 teeth, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, TOWEL CLAMP, BACKAUS, 13 cm | 6
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | RETRACTOR, VOLKMANN, 22 cm, 3sharp prongs, 10 mm curve | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | RETRACTOR, VOLKMANN, 22 cm, 4blunt prongs, 10 mm curve | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | RETRACTOR, LANGENBECK, 21 cm,16 x 28 mm | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | RETRACTOR, SELF-RET., WEITLANER, 16 cm, 3 x 4 blunt prongs | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | RETRAC., SELF-RET., ADSON, 31cm, 4 x 5 blunt pr., laminect. | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | FORCEPS, TISSUE, ALLIS, 15 cm/4x5 teeth, standard | 2
SET, GENERAL SURGERY INSTRUMENTS, BASIC SURGERY | BASKET, STERILIZING, 240x300x50mm,with cover and handles | 1
SET, GENERAL SURGERY INSTRUMENTS, DRESSING | KIDNEY DISH, medium, 250x140x40mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, DRESSING | SCISSORS, 14.5 cm, sharp-blunt, straight | 1
SET, GENERAL SURGERY INSTRUMENTS, DRESSING | FORCEPS, DRESSING, BLANK, 14.5 cm, atraumatic serration | 1
SET, GENERAL SURGERY INSTRUMENTS, DRESSING | FORCEPS, PEAN, 14 cm, serrated jaws, half way | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | FORCEPS, ADSON, 12 cm, serrated jaws | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | FORCEPS, ADSON, 12 cm, 1x2 teeth | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | SUCTION TUBE, FRAZIER, CH 12,12.5 cm, angular | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | FORCEPS, HEMOSTATIC, H-MOSQUITO, 12.5 cm, curved | 10
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | FORCEPS, HEMOSTATIC, DANDY, 14 cm, serrated, sideways curved | 8
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | RETRACTOR, HOOK, SKIN, BARSKY,15 cm, single prong, sharp | 2
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | RASPATORY, FARABEUF, 150x13mm, light curved, sharp | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | RONGEUR, BONE, FRIEDMANN, 14 cm x 2,5 mm, curved, delicate | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | FORCEPS, BONE CUT., LISTON, 17 cm, standard, straight | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | SCISSORS, VASCULAR, POTTS-SMITH, 19cm, 60deg angled | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | RONGEUR, LAMINECTOMY, SPURLING-KERRISON, 20cm, jaws 5mm | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | DISSECTOR, DURA, OLIVECRONA, 18 cm, double ended 4 + 5 mm | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | DRILL, HAND, CRANIAL, HUDSON,complete set | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | (drill, hand, cranial, Hudson) DRILL, CUSHING, 14mm, flat | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | (drill, hand, cranial, Hudson) PERFORATOR, McKENSIE, 13mm | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | (drill, hand, cranial, Hudson) | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | SAW GIGLI, HANDLE, inpairs | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | SAW GIGLI, WIRE, 50 cm | 5
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | SAW GIGLI, GUIDE, DeMARTEL, 33 cm | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | CANNULA, TEAR DUCT, RANDOLPH,15mm, Luer-Lock | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | LEVER, BONE, LANE, 21 cm | 1
SET, GENERAL SURGERY INSTRUMENTS, CRANIOTOMY, complementary | BASKET, STERILIZING, 240 x 255 x 100mm + lid 24x25cm | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | BASKET, STERILIZING, 480 x 255 x 100mm,Ref 49.97.13 + lid 48x25cm | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | CLAMP, VASCULAR, AORTA, DeBAKEY, 18 cm, pronouncedly shaped | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, DRESSING, BLANK, 14.5 cm, atraumatic serration | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, DRESSING, STANDARD, 20 cm, straight | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 10
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, HEISS, 24 cm, curved, right-angle | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, ROCHESTER-PEAN, 20 cm, curved | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, ROCHESTER-PEAN, 26 cm, curved | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, R-OCHSNER, 16 cm/1x2 teeth, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HEMOSTATIC, R-OCHSNER, 22 cm/1x2 teeth, straight | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, HYSTERECTOMY, WERTHEIM, 23 cm, 1 x 2 teeth, curved | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, INTESTINAL, KOCHER, 26cm, elast. atraum., straight | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, INTESTINAL, KOCHER, 26cm, elast. atraum., curved | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, TISSUE, ALLIS, 15 cm/4x5 teeth, standard | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, TISSUE, BABCOCK, 16 cm, 9 mm jaws, standard | 4
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, TISSUE, LANE, 14cm, 1x2 teeth | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, TISSUE, STANDARD, 20cm, 1x2 teeth, straight | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS, Towel Clamp, Balls and Socket, Backhaus | 8
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | FORCEPS-CLAMP, UTERINE, GREEN-ARMYTAGE, 22 cm x 13 mm | 6
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | HOLDER, INSTRUMENTS, BUNT (MAYO), 14 cm | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | KIDNEY DISH, medium, 250x140x40mm, stainless steel | 2
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | NEEDLE HOLDER, MAYO-HEGAR, 15cm, standard | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | NEEDLE HOLDER, MAYO-HEGAR, 18cm, standard | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | NEEDLE HOLDER, MAYO-HEGAR, 23cm, standard | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | PHOTOMACROGRAPHIC SCALE, 100mm, black aluminum, PPSM300 | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | RETRACTOR, ABDOMINAL, BALFOUR, standard, complete | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | RETRACTOR, FRISCH, 25.5 cm, 75 x 45 mm | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | RETRACTOR, ROUX, complete set(sizes 1, 2, 3) | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | RETRACTOR/SPATULA, DAEVER, 30cm x100mm, flexible | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SCALPEL, HANDLE, No 3 (for blades 10/11/15) | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SCALPEL, HANDLE, No 4 (for blades 20) | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SCISSORS, MAYO, 17 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SCISSORS, METZENBAUM, 14 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SCISSORS, METZENBAUM, 20 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, LAPAROTOMY, + caesarian | SUCTION TUBE, POOLE, 23 cm/diam.10mm, straight | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | KIDNEY DISH, large, 275x150x45mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | SCALPEL, HANDLE, No 3 (for blades 10/11/15) | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | SCISSORS, MAYO, 17 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | SCISSORS, METZENBAUM, 18 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | FORCEPS, TISSUE, LANE, 14cm, 1x2 teeth | 2
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | NEEDLE HOLDER, MAYO-HEGAR, 15cm, standard | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | NEEDLE HOLDER, MAYO-HEGAR, 18cm | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 4
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | FORCEPS, HEMOSTATIC, KOCHER, 14 cm/1x2 teeth, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | FORCEPS, Towel Clamp, Balls and Socket, Backhaus | 6
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | CURETTE, BONE, VOLKMANN, 22 cm, double-ended, fig. 3 and 5 | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, DPC (suture) | BASKET, STERILIZING, 254 x 410 x 70 mm+ lid 24x25cm | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | Photomacrographic autoclavable scale, right angle | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | KIDNEY DISH, large, 275x150x45mm, stainless steel | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | SCALPEL, HANDLE, No 3 (for blades 10/11/15) | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | SCALPEL, HANDLE, No 4 (for blades 20) | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | SCISSORS, MAYO, 17 cm, curved | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | SCISSORS, METZENBAUM, 18 cm, curved | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | SCISSORS, METZENBAUM, 20 cm, curved | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, TISSUE, LANE, 14cm, 1x2 teeth | 2
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | NEEDLE HOLDER, MAYO-HEGAR, 15cm, standard | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 8
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, HEMOSTATIC, R-OCHSNER, 16 cm/1x2 teeth, straight | 4
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, Towel Clamp, Balls and Socket, Backhaus | 6
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | RETRACTOR, VOLKMANN, 22 cm, 3sharp prongs, 10 mm curve | 2
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | RETRACTOR, LANGENBECK, 21 cm,13 x 42 mm | 2
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | RETRACTOR, SELF-RET., WEITLANER, 16 cm, 3 x 4 blunt prongs | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | RETRAC., SELF-RET., ADSON, 31cm, 4 x 5 blunt pr., laminect. | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | CURETTE, BONE, VOLKMANN, 22 cm, double-ended, fig. 3 and 5 | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | RONGEUR, BONE, LUER, 18 cm x 10 mm, heavy pattern | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | FORCEPS, TISSUE, ALLIS, 15 cm/4x5 teeth, standard | 2
SET,GENERAL SURGERY INSTRUMENTS, DEBRIDEMENT | BASKET, STERILIZING, 480 x 240 x 90 mm + lid | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | PHOTOMACROGRAPHIC SCALE, 100mm, black aluminum, PPSM300 | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | SCISSORS, METZENBAUM, 14 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | FORCEPS, DRESSING, BLANK, 14.5 cm, atraumatic serration | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | FORCEPS, Towel Clamp, Balls and Socket, Backhaus | 6
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | BOWL, ROUND, 200ml, 9.5 x 5 cm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | DERMATOME, HUMBY | 1
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | (dermatome Humby), BLADE | 5
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | (dermatome Schink) PLATE, forskin holding | 2
SET, GENERAL SURGERY INSTRUMENTS, SKIN GRAFT | BASKET, STERILIZING, 240 x 255 x 100mm, + lid 24x25cm | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | PHOTOMACROGRAPHIC SCALE, 100mm, black aluminum, PPSM300 | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | SCALPEL, HANDLE, No 3 (for blades 10/11/15) | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | SCISSORS, MAYO, 17 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | SCISSORS, METZENBAUM, 14 cm, curved | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, ADSON, 12 cm, serrated jaws | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, ADSON, 12 cm, 1x2 teeth | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, McINDOE, 15 cm, serrated | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, GILLIES, 15 cm, 1x2 teeth | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | PROBE, 16 cm/diam. 1.5 mm, round double-ended | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | SUCTION TUBE, FRAZIER, CH 12,12.5 cm, angular | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | NEEDLE HOLDER, BABY-CRILE-WOOD, 15 cm | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, HEMOSTATIC, H-MOSQUITO, 12.5 cm, curved | 10
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 4
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, HEMOSTATIC, KOCHER, 14 cm/1x2 teeth, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, TOWEL CLAMP, BACKAUS, 13 cm | 6
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | RETRACTOR, SENN-MILLER, 16 cm, double-ended, blunt prong | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | RETRACTOR, VOLKMANN, 22 cm, 3sharp prongs, 10 mm curve | 2
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | RETRACTOR, SELF-RET., WEITLANER, 13 cm, 3 x 4 blunt prongs | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, TISSUE, ALLIS-BABY, 14 cm / 3x4 theeth | 2
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | FORCEPS, TISSUE, BABCOCK, 16 cm, 9 mm jaws, standard | 2
SET, GENERAL SURGERY INSTRUMENTS, FINE(pediat),complementary | BASKET, STERILIZING, 240 x 255 x 100mm, + lid 24x25cm | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | FORCEPS, TISSUE, DUVAL, 23cm,atraum., triangular | 2
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | RASPATORY, FARABEUF, 150x13mm, light curved, sharp | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | MALLET, 23 cm x 28 mm, 490 g,solid | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | SHEAR, RIB, GIERTZ, 25cm | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | CHISEL, STERNUM, LEBSCHE, 25cm | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | SPREADER, RIB, FINOCCHIETTO, medium, blade, 65x45mm | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | CONTRACTOR, RIB, BAILEY, 17cm | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | RASPATORY, RIB, DOYEN, 17cm, right | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | RASPATORY, RIB, DOYEN, 17cm, left | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | RETRACTOR, LUNG, ALLISON, 270x40mm | 2
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | CLAMP, VASCULAR, DeBAKEY, 22 cm, straight | 2
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | CLAMP, VASCULAR, AORTA, DeBAKE | 1
SET, GENERAL SURGERY INSTRUMENTS, THORACOTOMY, complementary | BASKET, STERILIZING, 480 x 255 x 100mm,+ lid 48x25cm | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | BIT, DRILL, 3.2 mm.Jacobs chuck, 180/165 mm | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | PHOTOMACROGRAPHIC SCALE, 100mm, black aluminum, PPSM300 | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | KIDNEY DISH, medium, 250x140x40mm, stainless steel | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | SCALPEL, HANDLE, No 4 (for blades 20) | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | SCISSORS, MAYO, 15.5 cm, straight | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | SCISSORS, METZENBAUM, 18 cm, curved | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, DRESSING, BLANK, 14.5cm, atraumatic serration | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, TISSUE, LANE, 14cm, 1x2 teeth | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | NEEDLE HOLDER, MAYO-HEGAR, 15cm, standard | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, HEMOSTATIC, CRILE, 14 cm, curved | 6
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, HEMOSTATIC, R-OCHSNER, 16 cm/1x2 teeth, straight | 4
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, Towel Clamp, Balls and Socket, Backhaus | 6
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, SPONGE, FOERSTER, 24cm, serrated jaws, straight | 2
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | RETRACTOR, VOLKMANN, 22 cm, 1sharp prong, 20 mm curve | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | RETRACTOR, VOLKMANN, 22 cm, 3sharp prongs, 10 mm curve | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FILE, BONE, 22 cm x 20 mm, flat | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | RASPATORY, FARABEUF, 150x13mm, light curved, sharp | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | CURETTE, BONE, VOLKMANN, 22 cm, double-ended, fig. 3 and 5 | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | RONGEUR, BONE, LUER, 18 cm x 10 mm, heavy pattern | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, BONE CUT., LISTON, 17 cm, standard, straight | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, BONE CUT., LISTON-STILLE, 27 cm, dou. act., straigh | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | BOWL, ROUND, 100 ml, 80 x 35 mm, stainless steel | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | FORCEPS, TISSUE, ALLIS, 15 cm/4x5 teeth, standard | 2
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | SAW GIGLI, HANDLE, inpairs | 1
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | SAW GIGLI, WIRE, 50 cm | 10
SET, ORTHO. SURGERY INSTRUMENTS, AMPUTATION | BASKET, STERILIZING, 480 x 255 x 100mm,+ lid 48x25cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | PHOTOMACROGRAPHIC SCALE, 100mm, black aluminum, PPSM300 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | BONE HOOK, SHARP, VOLKMANN, 15 mm, 21cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RETRACTOR, VOLKMANN, 22 cm, 1sharp prong, 20 mm curve | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | FILE, BONE, 22 cm x 20 mm, flat | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RASPATORY, FARABEUF, 150x13mm, light curved, sharp | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | MALLET, 23 cm x 28 mm, 490 g,solid | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | OSTEOTOME, STILLE, 20 cm x 15mm, straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | OSTEOTOME, STILLE, 20 cm x 20mm, straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | GOUGE, STILLE, 20 cm x 15 mm,straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | GOUGE, STILLE, 20 cm x 20 mm,straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | CURETTE, BONE, VOLKMANN, 22 cm, double-ended, fig. 2 and 4 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | CURETTE, BONE, VOLKMANN, 22 cm, double-ended, fig. 3 and 5 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RONGEUR, BONE, LUER, 15 cm x 5 mm, strong curve | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RONGEUR, BONE, BANE, 18 cm x 5 mm, curved | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RONGEUR, BONE, LUER, 18 cm x 10 mm, heavy pattern | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RONGEUR, BONE, STILLE- LUER, 23 cm x 10 mm, double action | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | FORCEPS, BONE CUT., LISTON, 17cm, standard, straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | FORCEPS, BONE CUT., LISTON-STILLE, 27 cm, dou. act., angled | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | FORCEPS, BONE HOLDING, FERGUSON, 21 cm/4 x 4 teeth | 2
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | FORCEPS, BONE HOLD., VERBRUGGE, 27 cm x 12 mm, dismountable | 2
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | RETRACTOR, BONE, HOHMANN, 29 cm x 34 mm | 2
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG.,complement. | BASKET, STERILIZING, 240 x 255 x 100mm, + lid 24x25cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN, 17 cm, fig. 0000 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN, 17 cm, fig. 00 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN, 17 cm, fig. 2 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN, 17 cm, fig. 4 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN, 17 cm, fig. 6 | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN-Bruns, 23cm, fig 2, oval, straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | CURETTE, BONE, VOLKMANN-Bruns, 23cm, fig 5, oval, straight | 1
SET, ORTHO.SURGERY INSTRUMENTS, BASIC BONE SURG., curettes | BASKET, STERILIZING, 275 x 125 x 40mm, + lid 27,5x12,5 cm | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | BASKET, STERILIZING, 272 x 172 x 70mm + lid | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | CHUCK, UNIVERSAL, (3 jaws) with Thandle + locking device | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | PLIERS, FLAT NOSE, 14cm, serrated jaws | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | PLIERS, FLAT NOSE, 16 cm, serrated jaws, heavy pattern | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | PLIERS, WIRE CUTTING, 26cm, hard wire up to diam. 3.2mm | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | VICE GRIP, 18 cm, self locking | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE PASSER, DEMEL, diam. 45 mm x 28 cm, blunt | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE PASSER, diam 45mm x 23cm, curved to left, blunt | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, CERCLAGE, 1.00 mm x 10 m, stainless steel, coil | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, CERCLAGE, 1.25 mm x 10 m, stainless steel, coil | 1
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, KIRSCHNER, Ø 1.00mm x 15cm, trocard + round end | 10
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, KIRSCHNER, Ø 1.60mm x 15cm, trocard + round end | 10
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, KIRSCHNER, Ø 2.00mm x 15cm, trocard + round end | 10
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, KIRSCHNER, Ø 2.50mm x 15cm, trocard + round end | 10
SET, ORTHO. SURGERY INSTRUMENTS, BONE WIRING and KIRSHNER | WIRE, KIRSCHNER, Ø 3.00mm x 15cm, trocard + round end | 10
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | SCISSORS, PLASTER SHEARS, BRUNS, 24cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | SCISSORS, PLASTER SHEARS, STILLE-mini, 20cm, for small POP | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | SCISSORS, PLASTER SHEARS, STILLE, 37 cm, reinforced blade | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | KNIFE, PLASTER, ESMACH, 18 cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | SPREADER, PLASTER, HENNIG, 27cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | BREAKER, PLASTER CAST, WOLFF-BOEHLER, 18 cm | 1
SET, ORTHO.SURGERY INSTRUMENTS, PLASTER CASTS REMOVAL | BASKET, STERILIZING, 254 x 410 x 70 mm | 1
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | PIN, STEINMANN, 4.0 mm x 150 mm | 3
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | BIT, DRILL, 3.2 mm.Jacobs chuck, 180/165 mm | 1
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | CHUCK, UNIVERSAL (3-jaw) withT-handle + locking device | 1
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | ROPE, TRACTION, 3mm x 15m, bow, extension, Boehler | 2
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | SPLINT, traction frame | 5
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | BOW, EXTENSION, BOEHLER, 9 x 16 cm | 2
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | BOW, EXTENSION, BOEHLER, 15x21cm | 8
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | NAIL, EXTENSION, STEINMANN, 210 x 4 mm, trocar point | 7
SET, ORTHO.SURGERY INSTRUMENTS, TRACTION, + 10 bows | BASKET, STERILIZING, 480 x 240 x 90 mm + lid | 1